# Drawing template

Place your acrylic electron template face up on the table. Measure two principal axis and record the number. This will be used to compare in Raystation later.

Place your clear plastic sheet on top and trace cutout. Note where -x and -y are

Place the traced plastic face down on the scanner, with -X and -Y in the same direction as shown in the image below.

Cover it with a piece of white paper

## Scan the image

Preview the scan to make sure it looks okay

## Set resolution

Make sure to set resolution to 72 DPI

## Save the image

Save the file as a .jpg image, uploading it to

\\vscifs1\physicsQAdata\UNC_ElectronCutout\*marker color*

After uploading, a mask file will automatically be generated, review it now!

## Review image

A jpeg file called ‘mask’ will be generated, make sure it looks acceptable.

Delete all but the most recently created ‘RS_MRN*.dcm’ files!

# Raystation Importing

Open Electron phantom

## Open the Template Plan

## Copy Plan

Make a copy of this plan by going to Plan design -> Copy plan

Name plan with First_Last_MM.DD.YYYY format

## Set Beam SSD

Beam SSD To Surface should be set to 95.5. This means that aperture is almost exactly on the surface of the phantom

## Clear previous ROI of ‘New’

Select ‘Patient modeling’ -> ‘New’ Target and Delete Geometry

## Import generated structure from scanner

Once the .jpg file has been made, it will automatically create a structure file.

Click the blue button in the upper left corner -> Import -> Import to current patient

Select ‘File’ and paste

\\vscifs1\PhysicsQAdata\UNC_ElectronCutout

## Measure comparison

Localize ‘New’ and use the measure tool to evaluate the dimensions of the structure.

# Create cutout from structure

Go to Plan Design -> Electron beam design -> Cutout

## Verify beam SSD!

Beam SSD should still be 95.5!

## Set Applicator

Applicator defaults to be 25x25, set it to the desired size

## Conform Cutout

Select conform cutout, cutout should conform to the ‘New’ structure

## Change SSD

Change the SSD to the desired SSD for treatment